Soluble Reactive Phosphorus Protocol

Instrument Specifications - Hach DR1900

- Method 8048: USEPA PhosVer 3 (Ascorbic Acid) Method

- Range: 0.02 - 2.50 mg/L PO43-

- Precision (95% Confidence Interval) for 2.00 mg/L standard: 1.98 - 2.02 mg/L PO43-

- Sensitivity Concentration change per 0.010 Abs change: 0.02 mg/L PO43-

Equipment and Supplies

- DR1900 portable spectrophotometer (4 AA batteries)

- 2 to 4 sample cells for 10 mL volume (Product #2495402)

- PhosVer 3 Phosphate Reagent Powder Pillow, 10-mL

- 50.0 mg/L Phosphate Standard Solution (PO4) → 500 mL

- 1 volumetric flask (Class A, 100 mL capacity)

- Pipette, 1.0 mL with tips

- Stoppers for 18 mm tube

- Transfer pipets

- Gloves

- Kimwipes

- Thermometer

- Nutrient Analysis Data Sheet

- 1N HCl

- Deionized water

- 2 clean squeeze bottles (can be the same as what is prepped for Silica)

- 250-mL dark amber plastic Nalgene bottles

Important Notes and Considerations

- Clean sample bottles and sample cells with 1N HCl and rinse with deionized water before sample collection and testing.

- While wearing gloves, collect samples in clean 250 mL dark amber plastic Nalgene bottles.

- Ensure the samples have been filtered before analysis.

- Analyze the samples soon after collection for best results.

  - If prompt analysis is not possible, keep the samples in the dark at or below 6°C for a maximum of 48 hours.

- Samples should be at room temperature (~19 - 22°C) when being analyzed.

- Ensure the lab is at a stable temperature (± 2°C), check the first and last sample temperatures during analysis.

General Preparation

- Follow the ‘Water Sampling_SOP’ to collect water samples at specified depths and fill the clean 250-mL dark amber plastic bottles.

- Follow the ‘Filtering a Water Sample_SOP’ to filter the sample water either on the boat/dock or back in the lab. 10 mL of water is needed for a single phosphate test but an additional 10 mL is needed for a blank measurement of the same water.

  - Filtered samples can be transferred into a clean bottle for analysis.

- Print the Nutrient Analysis Data Sheet to record all unique Sample IDs, nutrient concentration results, and to make any comments during analysis.

- If users know which depths and stations will be routinely sampled, the user can make unique sample IDs that are saved on the DR1900 nutrient analyzer ahead of the analysis day. The sample IDs can be accessed easily vs. inputting and making new IDs during analysis.

- Individual users can also create and save their names as an ‘Operator ID’ to the instrument which will be added to the sample datafile.

- Label the benchtop/workspace in front of each sample cell (if more than 2 are accessible) with either lab tape or a sticky note to avoid any mix-ups.

  - e.g.: Blank, Standard, Sample 1, Sample 2.

- Record the lab temperature at the start of the analysis day and record the first and last sample temperatures onto the Nutrient Analysis Data Sheet to assess temperature change during the analysis.

- Always wear gloves when analyzing the nutrient samples.

Soluble Reactive Phosphorus Analysis

- Standard Preparation and Reaction Steps:

  - Using the 1 mL pipette, add 4 x 1.0 mL additions (total 4.0 mL) of 50 mg/L Phosphate Standard Solution to the 100 mL volumetric flask.

    - Users can gravimetrically check that the 1 mL pipette is accurately pipetting by placing the volumetric flask on a balance and adding the standard additions up to 4 mL total.

  - Add 96 mL of deionized water to the volumetric flask to dilute the standard solution to 2.0 mg/L.

    - Fill to the calibration mark on the volumetric flask ensuring the meniscus is on the mark.

  - Put the glass stopper on the volumetric flask and mix the standard solution to homogenize.

  - Using a new transfer pipet, add 10 mL of the freshly prepared 2.0 mg/L Phosphate Standard Solution into a clean sample cell. Fill the sample cell to the 10 mL mark ensuring the meniscus of standard solution is on the mark.

  - Open one PhosVer 3 Phosphate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the standard sample cell. Tap/flick the packet to ensure all the PhosVer 3 reagent has been added to the sample cell.

  - Put a rubber stopper into the sample cell opening and shake the sample cell for 30 seconds to dissolve the PhosVer 3 reagent.

  - Start a timer for a 2 minute reaction.

    - A blue colour indicates the presence of phosphorus.

    - While the 2 minute reaction time is happening, users can prepare the blank sample cell.

- Sample Preparation and Reactions Steps:

  - Insert the thermometer into the first and last sample and record the temperature onto the analysis data sheet (it should be at room temperature between 20° to 22°C, depending on lab conditions).

  - If users want to perform a Biological Replicate, the same sample water will need to be transferred into three different sample cells and undergo the preparation and reaction steps laid out in this SOP.

    - The Biological Replicate doesn’t need to be performed during every analysis, but it is good practice to do occasionally.

  - Using a new transfer pipet, add 10 mL of sample water into a clean sample cell. Fill the sample cell to the 10 mL mark ensuring the meniscus of sample water is on the mark.

  - Open one PhosVer 3 Phosphate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the sample water sample cell. Tap/flick the packet to ensure all the PhosVer 3 reagent has been added to the sample cell.

  - Put a rubber stopper into the sample cell opening and shake the sample cell for 30 seconds to dissolve the PhosVer 3 reagent.

  - Start a timer for a 2 minute reaction.

    - A blue colour indicates the presence of phosphorus.

    - If the blank measurement has not been done yet, prepare the blank sample cell.

- Blank Preparation Steps:

  - During the 2 minute reaction time, prepare the blank cell.

  - Using a new transfer pipet, add 10 mL of the same sample water into a clean sample cell. Fill the sample cell to the 10 mL mark ensuring the meniscus of sample water is on the mark.

Measuring Soluble Reactive Phosphorus

- The DR1900 instrument should be turned on by pressing the Power key.

  - The DR1900 always undergoes a ‘Self-Check’ upon start-up:

    - System Tests; Lamp Tests; Voltage Tests; Wavelength Calibrations will all be checked.

- Push the < SETTINGS > button on the DR1900 instrument (the wrench symbol).

- Select the ‘All Programs / Methods’ option in the menu.

- Select the ‘Stored Programs / Methods’ option in the menu.

- Select the ‘All Programs / Methods’ option in the dropdown menu.

- Scroll to find the program ‘490 P React. PV 2.50 mg/L’, press ‘Start’ for the program to load.

  - After starting the program, the instrument will automatically go back to the main screen and the selected program name and number will be displayed in the top left corner of the screen.

  - The appropriate wavelength for measurement (710 nm) will also be shown in the top right corner next to the battery status bar.

- Push ‘Options’ on the DR1900 instrument to access the Option Menu.

  - Users will select their appropriate ‘Operator ID’. (A maximum of 50 names can be entered).

  - Unique ‘Sample IDs’ can be selected to associate the readings with a particular sample and/or location. (A maximum of 12 characters can be entered).

    - They can be made beforehand (recommended) or created before the blank/standard/sample is read on the instrument. (A maximum of 100 names can be entered).

- Write down all unique ‘Sample IDs’ and results on the Nutrient Analysis Data Sheet along with any comments that may occur when measuring the blank/standard/sample(s).

- Blank Reading:

  - Select the unique ‘Sample ID’ for the ‘blank’ to be measured and the system to be zeroed (e.g., SRP_Blank).

  - Clean the exterior of the blank sample cell with a kimwipe to remove any smudge marks or water droplets and insert the blank sample cell into the cell holder.

    - The 10 mL mark should be facing towards the right when placed inside the holder.

  - Close the sample cell cover to ensure no light interferes with the measurement.

  - Push < ZERO >, the display will show 0.00 mg/L PO43-.

  - Remove the blank sample cell from the cell holder.

    - Set aside if the blank will be read as a “sample” later on during analysis if more than 5 samples will be analyzed in a day.

    - Reading the blank as a “sample” ensures that no optical drift has occurred over the analysis day, i.e., the value should still read 0.00 mg/L PO43-.

- Standard Reading:

  - Once the 2 minute reaction timer has expired, the ‘standard’ is ready to be measured.

  - Select the unique ‘Sample ID’ for the ‘standard’ to be measured (e.g. SRP_2MGL_STD).

  - Clean the exterior of the standard sample cell with a kimwipe to remove any smudge marks or water droplets and insert the standard sample cell into the cell holder.

    - The 10 mL mark should be facing towards the right when placed inside the holder.

  - Close the sample cell cover to ensure no light interferes with the measurement.

  - Push < READ >, the display will show mg/L PO43-.

    - The standard should read between 1.98 to 2.02 mg/L for confidence in reaction chemistry and DR1900 instrument response. If the standard reads below this value, see the Troubleshooting section.

  - To perform Technical Replicate, read the standard sample cell two more times and record the values of each reading.

    - The Technical Replicate doesn’t need to be performed during every analysis, but it is good practice to do occasionally.

    - The ‘Sample ID’ can remain the same and will be numbered in sequence for each measurement until a different ID is selected. (e.g., SRP_2MGL_STD (001), SRP_2MGL_STD (002), SRP_2MGL_STD (003)).

  - Remove the vial from the cell holder.

- Sample Reading:

  - Once the 2 minute reaction timer has expired, the ‘sample’ is ready to be measured.

  - Select the unique ‘Sample ID’ for the ‘sample’ to be measured (e.g. SRP_0M_DELTA).

  - Clean the exterior of the sample water sample cell with a kimwipe to remove any smudge marks or water droplets and insert the sample water sample cell into the cell holder.

    - The 10 mL mark should be facing towards the right when placed inside the holder.

  - Close the sample cell cover to ensure no light interferes with the measurement.

  - Push < READ >, the display will show mg/L PO43-.

  - Remove the sample cell from the cell holder.

  - If a Biological Replicate was performed, the ‘Sample ID’ will remain the same for all replicate samples and the IDs will be affixed with (001, 002, 003, etc.) and should be recorded on the analysis data sheet.

- Sample Cell Cleaning:

  - Sample cells will need to be cleaned after blank/standard/sample measurements if:

    - More samples need to be analyzed for Soluble Reactive Phosphorus

    - Samples still need to be analyzed for Total Silica

    - All analysis is completed and final clean up is happening.

  - Rinse the sample cells 3 times with 1N HCl.

  - Rinse the sample cells 3 times with deionized water.

  - If the sample cells won’t be used after the cleaning and rinsing steps, prop the sample cells so they can air dry.

- If the next sample is added to these freshly cleaned and rinsed sample cells, add ~ 1 mL of sample, 3 times, to rinse and homogenize the sample cell with the sample water that will be measured and to minimize any carry-over effects.

- Safely dispose of the solutions in the sample cells according to regional waste management guidelines.

Data Download and Management

- Sample readings are saved automatically each time a sample is measured and can store up to 500 data points on the DR1900 instrument.

  - When the data log becomes full, the oldest data point is erased when a new data point is added.

  - It is recommended that users download data soon after the analysis is complete to ensure no data is lost.

- Make sure the DR1900 is turned off before installing the USB+Power module.

- Remove the connector cover piece and insert the USB+Power module into the slot on the underside of the DR1900.

  - Note: Do not connect the USB+Power module to the computer when the module is still connected to the DR1900 instrument.

- Push the < SETTINGS > button on the DR1900 instrument (the wrench symbol).

- Scroll down to select the ‘Data Management’ option.

- Scroll down to select the ‘Send Data’ option. All data that is stored on the DR1900 will be sent to the USB+Power module as a .TXT file.

  - Users can select the ‘Start Date’ option and enter in a specified date and time for collecting the data to be sent to the USB+Power module.

    - This option is useful if older data is stored on the DR1900 and the user only wants to look at a specific run date for the data download.

- The screen will read “Sending Data” while transferring data to the module and will display “Transfer Complete” when the data is successfully transferred to the module.

  - Exit the data transfer screen and turn off the DR1900 instrument.

- Remove the USB+Power module from the DR1900 instrument and reinstall the connector cover piece on the underside of the instrument.

- Using the USB to type A cable, plug the type A connector into the module and the USB connector into the computer.

- Open the DR1900 USB module on the computer.

  - It will show up on the computer as a USB storage device and will most likely have the name ‘2100Q USB’ associated with it.

- The data recently sent to the USB module may be found in the folder ‘datalog’ or in the main directory.

  - The filename from the DR1900 instrument will be formatted as: LOG-YYMMDDHHMMSS.TXT.

- Copy and paste the data file to a date specific folder on the computer and rename the file as: {YYYY-MM-DD}_nutrients_raw.txt.

- Open a copy of the {YYYY-MM-DD}_nutrients_raw.txt in excel and format the data using ‘text to columns’ with comma delimitation selected. Save the formatted file as: {YYYY-MM-DD}_nutrients_raw.csv.

- Make a copy of the ‘Nutrient Data Sheet Template’ and save as: {YYYY-MM-DD}_nutrients_processed.csv.

  - Fill in the ‘nutrients_processed’ sheet using information from the downloaded ‘nutrients_raw.csv’ file, the information from the ‘Nutrient Analysis Data Sheet’ and the information collected in the field from the ‘GEM Survey Log sheet’.

- Save a .csv version of the {YYYY-MM-DD}_nutrients_processed sheet and save as {YYYY-MM-DD}_nutrients_final.csv.

- Navigate to the /data/{YYYY-MM-DD}_survey/DR1900/ folder of the Github repository, and upload the following:

  - {YYYY-MM-DD}_nutrients_raw.txt

  - {YYYY-MM-DD}_nutrients_raw.csv

  - {YYYY-MM-DD}_nutrients_processed.csv

  - {YYYY-MM-DD}_nutrients_final.csv

- Take a photo of the ‘Nutrient Analysis Data Sheet(s)’ and upload to the same folder with the name {YYYY-MM-DD}_nutrients_datasheet_raw#.jpeg. Updated the # at the end if more than one datasheet was used during analysis.

Troubleshooting

- Problem: DR1900 doesn’t turn on.

  - Solution: Try changing the batteries.

  - Solution: If using the USB+Power module, ensure the power supply has a good connection on the USB+Power module and into the wall socket.

- Problem: Very low / high readings across sample replicates.

  - Solution: Ensure that there are no marks / smudges or liquid on the outside of the sample cell. These will affect the accuracy of the readings. Clean the outside surface of the sample cell with a kimwipe to remove and try reading the sample again.

  - Solution: Make sure the sample cell solutions do not contain any bubbles. Bubbles will affect the accuracy of the readings. Bubbles can be removed by gently tapping the outside of the sample cell wall.

- Problem: The DR1900 won’t allow the user to enter new ‘Operator IDs’ or ‘Sample IDs’ and proceeds to cycle through its power and start-up phase.

  - Solution: If the USB+Power module was inserted to download the data off the instrument, users will have to put the USB+Power module back into the instrument and plug the module into a wall outlet. This should resolve the problem for entering new ‘Operator IDs’ and ‘Sample IDs’.