Total Dissolved Inorganic Nitrogen Protocol

Instrument Specifications - Hach DR1900

- Method 10021: Titanium Trichloride Reduction Method

- Range: 0.2 - 25.0 mg/L N

- Precision (95% Confidence Interval) for 10.0 mg/L standard: 9.6 - 10.4 mg/L NO3-N

- Sensitivity Concentration change per 0.010 Abs change: 0.2 mg/L NO3-N

Equipment and Supplies

- DR1900 portable spectrophotometer (4 AA batteries)

- Adapter 9609900 (D1 - 16mm; included with DR1900 instrument)

- Protective cover (included with DR1900 instrument - # 9595500)

- Total Inorganic Nitrogen Pretreatment Reagent Set (TiCl3 Reduction Method)

- Test ’N Tube AmVer Nitrogen-Ammonia Reagent Set (Low Range Ammonia)

- 10.0 mg/L Nitrate Nitrogen Standard Solution (NO3-N) → 500 mL

- Centrifuge

- Pipette, 1.0 mL with tips

- Transfer pipets

- Test tube rack

- Water, deionized

- Gloves

- Kimwipes

- Funnel, micro

- Thermometer

- Nutrient Analysis Data Sheet

- Sharps container

- 250-mL dark amber plastic Nalgene bottles

Important Notes and Considerations

- While wearing gloves, collect samples in clean 250 mL dark amber plastic Nalgene bottles.

- Ensure the samples have been filtered before analysis.

- Analyze the samples soon after collection for best results.

  - If prompt analysis is not possible, keep the samples in the dark at or below 6°C for a maximum of 48 hours.

- Samples should be at room temperature (~19 - 22°C) when being analyzed.

- Ensure the lab is at a stable temperature (± 2°C), check the first and last sample temperatures during analysis.

- To ensure the 1 mL pipette is precisely adding the appropriate volume, gravimetrically check the volume by pipetting deionized water at room temperature into a weigh boat or beaker onto a balance.

  - Record the values of each mL addition and adjust the pipette if necessary to obtain ~ 1mL of solution.

- Upon receiving the centrifuge, users should set up the centrifuge to run for 3 minutes at 1000 rpm.

- Pipette tips should be re-used where possible during analysis and can be cleaned with 1N HCl and rinsed with deionized water for usage on the next analysis day(s).

General Preparation

- Follow the ‘Water Sampling_SOP’ to collect water samples at specified depths and fill the clean 250-mL dark amber plastic bottles.

- Follow the ‘Filtering a Water Sample_SOP’ to filter the sample water either on the boat/dock or back in the lab. 1 mL of water is needed for a single nitrogen test.

  - Filtered samples can be transferred into a clean bottle for analysis.

- Print the Nutrient Analysis Data Sheet to record all unique Sample IDs, nutrient concentration results, and to make any comments during analysis.

- If users know which depths and stations will be routinely sampled, the user can make unique sample IDs that are saved on the DR1900 nutrient analyzer ahead of the analysis day. The sample IDs can be accessed easily vs. inputting and making new IDs during analysis.

- Individual users can also create and save their names as an ‘Operator ID’ to the instrument which will be added to the sample datafile.

- The TIN Pretreatment Reagent Set vials and the Test ‘N Tube AmVer Reagent Set vials should be labeled with unique sample identifiers to avoid any mix-ups.

  - e.g: Blank 1A (for TIN Pretreatment Reagent vial) & Blank 1B (for AmVer Nitrogen-Ammonia Reagent vial).

  - e.g: Sample 1A & Sample 1B vs. Sample 2A & Sample 2B to denote two different samples that will be analyzed.

- Record the lab temperature at the start of the analysis day and only record the first and last sample temperatures onto the Nutrient Analysis Data Sheet to assess temperature change during the analysis.

- Always wear gloves when analyzing the nutrient samples.

Total Dissolved Inorganic Nitrogen Analysis

TIN Pretreatment Vial Preparation Steps:

- If users would prefer gravimetric additions vs volumetric additions, ensure that a clean / new pipette tip is used and a clean / new weigh boat for each solution that will be used in the following steps.

- Prepare the Blank Vial 1A:

  - Using a pipette and new tip, add 1 mL of Total Inorganic Nitrogen Pretreatment Base concentrate into the Total Inorganic Nitrogen Pretreatment Diluent ‘blank’ vial A.

  - Using a pipette and new tip, add 1 mL of deionized water to the Total Inorganic Nitrogen Pretreatment Diluent ‘blank’ vial A.

- Prepare the Standard Vial 1A:

  - Using a pipette and new tip, add 1 mL of Total Inorganic Nitrogen Pretreatment Base concentrate into the Total Inorganic Nitrogen Pretreatment Diluent ‘standard’ vial A.

  - Using a pipette and new tip, add 1 mL of 10.0 mg/L Nitrate Nitrogen Standard Solution to the Total Inorganic Nitrogen Pretreatment Diluent ‘standard’ vial A.

- Put the respective caps back onto each of the blank and standard vials and shake for 30 seconds to mix. Place the vials back onto the test tube rack.

- Prepare the Sample Vial #A:

  - Insert the thermometer into the first and last sample and record the temperature onto the analysis data sheet (it should be at room temperature between 20° to 22°C, depending on lab conditions).

  - If users want to perform a Biological Replicate, the same sample water will need to be pipetted into three different vials and undergo the preparation and reaction steps laid out in this SOP.

    - The Biological Replicate doesn’t need to be performed during every analysis, but it is good practice to do occasionally.

  - Using a pipette and new tip, add 1 mL of Total Inorganic Nitrogen Pretreatment Base concentrate into the Total Inorganic Nitrogen Pretreatment Diluent ‘sample’ vial A.

  - Using a pipette and new tip, add 1 mL of sample water to the Total Inorganic Nitrogen Pretreatment Diluent ‘sample’ vial A.

- Put the respective cap(s) back onto the vial(s) and shake for 30 seconds to mix. Place the vial(s) back onto the test tube rack.

TIN Pretreatment Vial Reaction Steps:

- Blank Vial 1A Reaction:

  - Remove the ‘blank’ vial cap and open one of the Total Inorganic Nitrogen Reductant ampules and use a transfer pipet to draw out the ampule contents.

  - Add the ampule contents to the ‘blank’ vial A.

    - Dispose of the empty ampule in a sharps container.

  - Cap the ‘blank’ vial and invert the solution once (↓ then ↑).

    - A black precipitate will form immediately after the ampule is added and if the vial is excessively shaken, the precipitate will turn gray/white.

  - Put the ‘blank’ vial back into the test tube rack and let it stand for ~ 90 seconds.

- Standard Vial 1A Reaction:

  - Remove the ‘standard’ vial cap and open one of the Total Inorganic Nitrogen Reductant ampules and use a transfer pipet to draw out the ampule contents.

  - Add the ampule contents to the ‘standard’ vial A.

    - Dispose of the empty ampule in a sharps container.

  - Cap the ‘standard’ vial and invert the solution once (↓ then ↑).

    - A black precipitate will form immediately after the ampule is added and if the vial is excessively shaken, the precipitate will turn gray/white.

  - Put the ‘standard’ vial back into the test tube rack and let it stand for ~ 90 seconds.

- Sample Vial #A Reaction:

  - Remove the ‘sample’ vial cap and open one of the Total Inorganic Nitrogen Reductant ampules and use a transfer pipet to draw out the ampule contents.

  - Add the ampule contents to the ‘sample’ vial A.

    - Dispose of the empty ampule in a sharps container.

  - Cap the ‘sample’ vial and invert the solution once (↓ then ↑).

    - A black precipitate will form immediately after the ampule is added and if the vial is excessively shaken, the precipitate will turn gray/white.

  - Put the ‘sample’ vial back into the test tube rack and let it stand for ~ 90 seconds.

- Place the vials into a centrifuge to spin for 3 minutes at 1000 rpm. The precipitate should be settled at the bottom of the vials.

  - Users can troubleshoot the speed of the centrifuge if a higher speed is better at keeping the precipitate down at the bottom of the vial.

  - If no centrifuge is available, let the vials stand for ~30 minutes or until the precipitate is settled at the bottom.

AmVer Nitrogen-Ammonia Vial Preparation Steps:

- After 3 minutes, remove the vials from the centrifuge and place in the test tube rack.

- Prepare the Blank Vial 1B:

  - Using a pipette and new tip, add 2 mL of the ‘Blank Vial 1A’ centrifuged solution into the AmVer Nitrogen-Ammonia Test ‘N Tube Diluent ‘blank’ vial B. Recap the vial.

  - Be careful not to disturb the precipitate that has settled at the bottom of the vial A.

  - Note: some bubbles will enter the pipette tip but they are negligible, ~2 mL of blank will still be transferred.

- Prepare the Standard Vial 1B:

  - Using a pipette and new tip, add 2 mL of the ‘Standard Vial 1A’ centrifuged solution into the AmVer Nitrogen-Ammonia Test ‘N Tube Diluent ‘standard’ vial B. Recap the vial.

  - Be careful not to disturb the precipitate that has settled at the bottom of the vial A.

  - Note: some bubbles will enter the pipette tip but they are negligible, ~2 mL of standard will still be transferred.

- Prepare the Sample Vial #B:

  - Using a pipette and new tip, add 2 mL of the ‘Sample Vial #A’ centrifuged solution into the AmVer Nitrogen-Ammonia Test ‘N Tube Diluent ‘sample’ vial B. Recap the vial.

  - Be careful not to disturb the precipitate that has settled at the bottom of the vial A.

  - Note: some bubbles will enter the pipette tip but they are negligible, ~2 mL of sample will still be transferred.

AmVer Nitrogen-Ammonia Vial Reaction Steps:

- Blank Vial 1B Reaction:

  - Optional to use a funnel for the next steps when adding the powder pillow reagents but note that reagent powder does get left on the funnels.

  - Uncap the ‘Blank Vial 1B’ and place the cap beside the test tube rack.

  - Open one Ammonia Salicylate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘blank’ vial B. Tap/flick the packet to ensure all the Salicylate reagent has been added to the vial.

  - Open one Ammonia Cyanurate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘blank’ vial B. Tap/flick the packet to ensure all the Cyanurate reagent has been added to the vial.

  - Put the cap back onto the ‘Blank Vial 1B’ and shake to dissolve the powder reagents completely.

  - Start a timer for a 20 minute reaction.

- Standard Vial 1B Reaction:

  - Optional to use a funnel for the next steps when adding the powder pillow reagents but note that reagent powder does get left on the funnels.

  - Uncap the Standard Vial 1B’ and place the cap beside the test tube rack.

  - Open one Ammonia Salicylate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘standard’ vial B. Tap/flick the packet to ensure all the Salicylate reagent has been added to the vial.

  - Open one Ammonia Cyanurate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘standard’ vial B. Tap/flick the packet to ensure all the Cyanurate reagent has been added to the vial.

  - Put the cap back onto the ‘Standard Vial 1B’ and shake to dissolve the powder reagents completely.

  - Start a timer for a 20 minute reaction.

    - A green colour indicates the presence of nitrogen.

- Sample Vial #B Reaction:

  - Optional to use a funnel for the next steps when adding the powder pillow reagents but note that reagent powder does get left on the funnels.

  - Uncap the Sample Vial #B’ and place the cap beside the test tube rack.

  - Open one Ammonia Salicylate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘sample’ vial B. Tap/flick the packet to ensure all the Salicylate reagent has been added to the vial.

  - Open one Ammonia Cyanurate Reagent Powder Pillow packet by tearing along the perforated line.

  - Make a U in the opening and add the contents into the ‘sample’ vial B. Tap/flick the packet to ensure all the Cyanurate reagent has been added to the vial.

  - Put the cap back onto the ‘Sample Vial #B’ and shake to dissolve the powder reagents completely.

  - Start a timer for a 20 minute reaction.

    - A green colour indicates the presence of Nitrogen.

Measuring Total Dissolved Inorganic Nitrogen

- Insert the ‘D1’ 16mm sample cell adapter into the sample cell location on the DR1900 instrument.

- Turn the DR1900 instrument on by pressing the Power key.

  - The DR1900 always undergoes a ‘Self-Check’ upon start-up:

    - System Tests; Lamp Tests; Voltage Tests; Wavelength Calibrations will all be checked.

- Push the < SETTINGS > button on the DR1900 instrument (the wrench symbol).

- Select the ‘All Programs / Methods’ option in the menu.

- Select the ‘Stored Programs / Methods’ option in the menu.

- Scroll to find the program ‘346 N Inorganic TNT 25.0 mg/L’, press ‘Start’ for the program to load.

  - After starting the program, the instrument will automatically go back to the main screen and the selected program name and number will be displayed in the top left corner of the screen.

  - The appropriate wavelength for measurement (655 nm) will also be shown in the top right corner next to the battery status bar.

- Push ‘Options’ on the DR1900 instrument to access the Option Menu.

  - Users will select their appropriate ‘Operator ID’. (A maximum of 50 names can be entered).

  - Unique ‘Sample IDs’ can be selected to associate the readings with a particular sample and/or location. (A maximum of 12 characters can be entered).

    - They can be made beforehand (recommended) or created before the blank/standard/sample is read on the instrument. (A maximum of 100 names can be entered).

- Write down all unique ‘Sample IDs’ and results on the Nutrient Analysis Data Sheet along with any comments that may occur when measuring the blank/standard/sample(s).

- Blank Vial 1B Reading:

  - Once the 20 minute reaction timer has expired, the ‘blank’ is ready to be measured.

  - Select the unique ‘Sample ID’ for the ‘blank’ to be measured and the system to be zeroed (e.g. DIN_Blank).

  - Clean the exterior of the ‘Blank Vial 1B’ with a kimwipe to remove any smudge marks or water droplets and insert the vial into the cell holder (ensure the D1 - 16mm cell adapter is in place).

  - Place the protective cover over the vial and sample cell housing to ensure no light interferes with the measurement.

  - Push < ZERO >, the display will show 0.0 mg/L N.

  - Remove the vial from the cell holder.

- Standard Vial 1B Reading:

  - Once the 20 minute reaction timer has expired, the ‘standard’ is ready to be measured.

  - Select the unique ‘Sample ID’ for the ‘standard’ to be measured (e.g. DIN_10MGL_STD).

  - Clean the exterior of the ‘Standard Vial 1B’ with a kimwipe to remove any smudge marks or water droplets and insert the vial into the cell holder.

  - Place the protective cover over the vial and sample cell housing to ensure no light interferes with the measurement.

  - Push < READ >, the display will show mg/L N.

    - The standard should read between 9.6 to 10.4 mg/L for confidence in reaction chemistry and DR1900 instrument response. If the standard reads below this value, see the Troubleshooting section.

  - To perform a Technical Replicate, read the standard vial two more times and record the values of each reading.

    - The Technical Replicate doesn’t need to be performed during every analysis, but it is good practice to do occasionally.

    - The ‘Sample ID’ can remain the same and will be numbered in sequence for each measurement until a different ID is selected. (e.g., DIN_10MGL_STD (001), DIN_10MGL_STD (002), DIN_10MGL_STD (003)).

  - Remove the vial from the cell holder.

- Sample Vial #B Reading:

  - Once the 20 minute reaction timer has expired, the ‘sample’ is ready to be measured.

  - Select the unique ‘Sample ID’ for the ‘sample’ to be measured (e.g., DIN_1M_DOCK).

  - Clean the exterior of the ‘Sample Vial #B’ with a kimwipe to remove any smudge marks or water droplets and insert the vial into the cell holder.

  - Place the protective cover over the vial and sample cell housing to ensure no light interferes with the measurement.

  - Push < READ >, the display will show mg/L N.

  - Remove the vial from the cell holder.

  - If a Biological Replicate was performed, the ‘Sample ID’ will remain the same for all replicate samples and the IDs will be affixed with (001, 002, 003, etc.) and should be recorded on the analysis data sheet.

- Safely dispose of the solutions in Vials A and B according to regional waste management guidelines. Dispose the sharps according to regional regulations.

Data Download and Management

- Sample readings are saved automatically each time a sample is measured and can store up to 500 data points on the DR1900 instrument.

  - When the data log becomes full, the oldest data point is erased when a new data point is added.

  - It is recommended that users download data soon after the analysis is complete to ensure no data is lost.

- Make sure the DR1900 is turned off before installing the USB+Power module.

- Remove the connector cover piece and insert the USB+Power module into the slot on the underside of the DR1900.

  - Note: Do not connect the USB+Power module to the computer when the module is still connected to the DR1900 instrument.

- Push the < SETTINGS > button on the DR1900 instrument (the wrench symbol).

- Scroll down to select the ‘Data Management’ option.

- Scroll down to select the ‘Send Data’ option. All data that is stored on the DR1900 will be sent to the USB+Power module as a .TXT file.

  - Users can select the ‘Start Date’ option and enter in a specified date and time for collecting the data to be sent to the USB+Power module.

    - This option is useful if older data is stored on the DR1900 and the user only wants to look at a specific run date for the data download.

- The screen will read “Sending Data” while transferring data to the module and will display “Transfer Complete” when the data is successfully transferred to the module.

  - Exit the data transfer screen and turn off the DR1900 instrument.

- Remove the USB+Power module from the DR1900 instrument and reinstall the connector cover piece on the underside of the instrument.

- Using the USB to type A cable, plug the type A connector into the module and the USB connector into the computer.

- Open the DR1900 USB module on the computer.

  - It will show up on the computer as a USB storage device and will most likely have the name ‘2100Q USB’ associated with it.

- The data recently sent to the USB module may be found in the folder ‘datalog’ or in the main directory.

  - The filename from the DR1900 instrument will be formatted as: LOG-YYMMDDHHMMSS.TXT.

- Copy and paste the data file to a date specific folder on the computer and rename the file as: {YYYY-MM-DD}_nutrients_raw.txt.

- Open a copy of the {YYYY-MM-DD}_nutrients_raw.txt in excel and format the data using ‘text to columns’ with comma delimitation selected. Save the formatted file as: {YYYY-MM-DD}_nutrients_raw.csv.

- Make a copy of the ‘Nutrient Data Sheet Template’ and save as: {YYYY-MM-DD}_nutrients_processed.csv.

  - Fill in the ‘nutrients_processed’ sheet using information from the downloaded ‘nutrients_raw.csv’ file, the information from the ‘Nutrient Analysis Data Sheet’ and the information collected in the field from the ‘GEM Survey Log sheet’.

- Save a .csv version of the {YYYY-MM-DD}_nutrients_processed sheet and save as {YYYY-MM-DD}_nutrients_final.csv.

- Navigate to the /data/{YYYY-MM-DD}_survey/DR1900/ folder of the Github repository, and upload the following:

  - {YYYY-MM-DD}_nutrients_raw.txt

  - {YYYY-MM-DD}_nutrients_raw.csv

  - {YYYY-MM-DD}_nutrients_processed.csv

  - {YYYY-MM-DD}_nutrients_final.csv

- Take a photo of the ‘Nutrient Analysis Data Sheet(s)’ and upload to the same folder with the name {YYYY-MM-DD}_nutrients_datasheet_raw#.jpeg. Updated the # at the end if more than one datasheet was used during analysis.

Troubleshooting

- Problem: DR1900 doesn’t turn on.

  - Solution: Try changing the batteries.

  - Solution: If using the USB+Power module, ensure the power supply has a good connection on the USB+Power module and into the wall socket.

- Problem: Very low / high readings across sample replicates.

  - Solution: Ensure that there are no marks / smudges or liquid on the outside of the vial. These will affect the accuracy of the readings. Clean the outside surface of the vial with a kimwipe to remove and try reading the sample again.

  - Solution: Make sure the sample cell or vial solutions do not contain any bubbles. Bubbles will affect the accuracy of the readings. Bubbles can be removed by gently tapping the outside of the vial.

- Problem: Precipitate in ‘TIN Pretreatment Vial Reaction Steps’ turns gray/white  with minimal mixing.

  - Solution: Check the Total Inorganic Nitrogen Reductant ampule expiration date to make sure it is not close to the expiration date, if it is, order new reagents.

- Problem: The DR1900 won’t allow the user to enter new ‘Operator IDs’ or ‘Sample IDs’ and proceeds to cycle through its power and start-up phase.

  - Solution: If the USB+Power module was inserted to download the data off the instrument, users will have to put the USB+Power module back into the instrument and plug the module into a wall outlet. This should resolve the problem for entering new ‘Operator IDs’ and ‘Sample IDs’.