Chlorophyll a and Turbidity Calibration and Performance Tracking Protocol

Equipment and supplies

- Turner AquaFluor handheld fluorometer/turbidimeter

- Turner Rhodamine Standard (400 ppb, 60 μg/L Chla equivalency, PN: 6500-120)

- AMCO 10 and 100 NTU Turbidity Standards (Thomas Scientific PN: C841E93/5)

- Turner Secondary Solid Standard (PN: 8000-952)

- Turner 12mm Test Tube Adapter (PN: 8000-932)

- VWR® 12mm Culture Tubes, Disposable, Borosilicate Glass (PN: 47729-570)

- 4 AAA batteries

- 5 mL Transfer pipettes

- KimWipes

- Nitrile gloves

- De-ionized (DI) water

Datasheets and Software

- AquaFluor Calibration Data Sheet

- AquaFluor Performance Data Sheet Template

Additional Resources

- Turner AquaFluor Manual: https://docs.turnerdesigns.com/t2/doc/manuals/998-0851.pdf

- Turner Liquid Dye Calibration Manual:

https://docs.turnerdesigns.com/t2/doc/tech-notes/S-0242.pdf

- Turner Solid Standard Configuration:

https://docs.turnerdesigns.com/t2/doc/instructions/998-0025.pdf

Notes

Handle culture tubes with gloves and clean with KimWipes. Carefully transfer standards and samples using transfer pipettes to limit cleaning and scratching of culture tubes. Dispose of excessively scratched tubes.

Fluorescence is temperature dependent. Calibrate in a temperature controlled room.

Use new culture tubes for calibration purposes.

Instrument Calibration

- Open the AquaFluor Calibration Data Sheet and fill in the calibration date and calibration number for the calibration.

- If possible, record the ambient temperature for the calibration in the AquaFluor Calibration Data Sheet.

- Place the 12mm culture tube adapter in the instrument sample compartment.

- Start with the IVCH channel - switch channels by pressing <A/B>.

  - The process is the same for the TRB channel with slight differences denoted in subpoints and bolded and italicized.

  - It is recommended to calibrate both channels every time a calibration is performed.

- Press the <CAL> button.

- Press <ENT>.

- Using a transfer pipette, ¾ fill a culture tube with DI water (blank).

- When prompted, place the DI water culture tube into the instrument sample compartment. Press <ENT>.

- Record the blank RFU value in the AquaFluor Calibration Data Sheet for the IVCH Channel.

- Press <ENT> (RFU no longer retrievable once pressed).

- Remove the blank culture tube.

- Use the <↑> /<↓> buttons to set the StdVal to 60.00 (Rhodamine Chl equivalent). Press and hold the buttons for faster scrolling. Press <ENT>.

  - TRB - 10.00 (NTU standard).

- Using a transfer pipette, ¾ fill a culture tube with Rhodamine standard.

  - TRB - 10 NTU Standard.

- When prompted, place the Rhodamine culture tube into the instrument sample compartment. Press <ENT>.

  - TRB - NTU Standard.

- Record the Rhodamine RFU value in the AquaFluor Calibration Data Sheet for the IVCH Channel under STD-1 (only one standard used for Chl).

  - TRB - TRB Channel under STD-1 (10 NTU) and STD-2 (100 NTU).

- Press <ENT> (RFU no longer retrievable once pressed).

- Use the <↑> /<↓> buttons to select End the calibration. Press <ENT>.

  - TRB - select measure another standard and repeat steps 10-15, but using the 100 NTU standard and recording the RFU under STD-2 in the AquaFluor Calibration Data Sheet.

- Use the <↑> /<↓> buttons to select μg/L. Press <ENT>. The screen should return to the main window.

  - TRB - NTU.

- Press the <STD VAL> button and make sure the Std Val is set to 1.0. Press <ESC>.

- Press <DIAG> and use the <↑> /<↓> buttons to select Cal Details.

- Record the Slope and fit in the AquaFluor Calibration Data Sheet under the appropriate columns

  - IVCH - fit will always be 1.000 with a single point calibration.

  - TRB - fit will vary and it is important to record.

- The instrument is now calibrated. This calibration is stored within the instrument until a new calibration is performed. The calibration will be automatically applied to field samples to derive outputs in μg/L Rhodamine Equivalent (RE).

  - TRB - NTU.

- Create a new Aquafluor Performance Data Sheet using the template. Save the sheet as {YYYY-MM-DD}_aquafluor_performance. Fill in the calibration date, number and temperature for the first row.

  - The calibration should always be placed in the first row. When a new calibration is performed, make a new sheet from the template.

- On the IVCH channel, remove the 12mm culture tube adapter and place the solid standard into the sample compartment with the tab facing the back of the instrument. Take a reading (press <READ>). Based on the result, adjust the solid standard and take additional readings until the solid standard shows a concentration close to 60 μg/L following: https://docs.turnerdesigns.com/t2/doc/instructions/998-0025.pdf. Record the finalized value in the top row of the SS Value column of the Aquafluor Performance Data Sheet.

  - The SS Change (%) column should remain empty.

- On the TRB channel, re-inset the 12mm culture tube adapter and place a culture tube  ¾ filled with 100 NTU standard into the sample compartment. Take a reading (press <READ>) and record it in the top row of the SS Value column of the Aquafluor Performance Data Sheet.

  - The SS Change (%) column should remain empty.

Calibration Tracking - Instrument Performance Monitoring

- Prior to running samples on each sample analysis day, fill in a new row of the Aquafluor Performance Data Sheet (use the calibration specific version saved as {YYYY-MM-DD}_aquafluor_performance).

  - The calibration date and number should remain the same for the entire sheet, but the analysis date will change.

  - Make a new sheet from the template to track a new calibration.

- Record the ambient temperature in the temp column.

- On the IVCH channel, place the solid standard into the sample compartment with the tab facing the back of the instrument. Take a reading (press <READ>) and record it SS Value column.

- On the TRB channel, place a culture tube  ¾ filled with 100 NTU standard into the sample compartment. Take a reading (press <READ>) and record it in the 100 NTU Value column.

- Once filled, drag the formula from the first empty rows (calibration) of the SS and NTU Change (%) columns downwards and the cells should populate.

  - If not, the % change formula is:

(measurement reading - calibration reading)/calibration reading)

- Check that both readings are within ± 10% of the calibration concentrations.

- If the readings are outside of the threshold:

  - IVCH - Check that the ambient temperature is comparable to the calibration temperature and run the solid standard again.

  - IVCH - Run the rhodamine standard to ensure that the deviation is due to the instrument and not the solid standard

  - TRB - Check that the culture tube  is clean and clear of bubbles. Try a new culture tube .

- If the above checks fail, recalibrate both channels of the instrument.